## COD_BOD_TSS_Temperature
Sl.No | Parameter | Configuration
1 | Model Number: | FLDBH6605D
2 | Output: | Rs-485,MODBUS/RTU
3 | Measurement Methods: | UV 254 Dual-wavelength ultraviolet absorption method
4 | COD Range: | 0～2000mg/L or Customize
5 | COD Accuracy: | ±5%F.S.
6 | COD Resolution: | 0.1mg/L or 0.01mg/L
7 | BOD Range: | 0~500mg/L or Customize
8 | TSS Range: | 0～1000mg/L
9 | Repeatability: | ±1%F.S.
10 | Working Condition: | 0.1～50℃、<0.3MPa
11 | Calibration: | Two points
12 | Response Time: | 10s T90
13 | Power Supply: | 12 or 24VDC±10%，15mA; 200mA when the brush rotates
14 | Dimensions: | Diameter 48 mm ; Length 186 mm; Optional protection cover
15 | IP Grade: | IP68，10m
16 | Flow Rate: | Less than 3m/s
17 | Life Time: | Sensor 3 years, cleaning brush system 18 months
18 | Cable Length: | 5m(default), customizable
19 | Sensor Material: | Titanium alloy, 316L, quartz glass
20 | Maintenance: | Sensor 3 months, cleaning brush 6 months
21 | Parameters: | COD, BOD, turbidity, and temperature
22 | Installation note | 1. Avoid collision when installing the sensor; avoid shaking back and forth during measurement. If the water flow in the monitoring area is turbulent, the sensor needs to be fixed; 2. When the sensor is installed and measured, it must be installed horizontally to prevent foreign matter from depositing on the measurement monitoring window, resulting in inaccurate measurement data; 3. The measurement area, such as branches and attachments, is easy to jam the cleaning brush shaft and cause the sensor to jam. You need to consider choosing a protective cover for protection installation; 4. Install the sensor, the depth from the water surface is no more than 2 meters; the sensor is submerged below the water level of 30cm; 5. The sensor is installed in an area with slow water flow and no bubbles
23 | Maintenance and precautions | 1. Measuring window: Check whether the measuring window is attached to and fouling; use a moist dust-free cloth to wipe the window; if the scale is difficult to wipe, use a dust-free cloth dampened with detergent to clean; or use diluted hydrochloric acid to wipe. Be careful not to touch it directly with your hands when using diluted hydrochloric acid. 2. Cleaning brush: Prevent foreign objects from blocking the rotation stroke of the cleaning brush, and ensure that no foreign objects are blocked at the cleaning rotating shaft; if there are many branches and attachments in the monitoring area, you need to consider selecting a protective cover for protection installation. The protective cover can be selected separately by contacting the sales staff of our company. 3. The outer surface of the sensor: Regularly clean the surface attachments, you can use washing liquid to clean 4. Automatic cleaning and continuous use for 18 months, need to return to the factory to replace the dynamic sealing device.
24 | Matters needing attention | 1. Measurement window: Sharp objects cannot be used to scratch the measurement monitoring window; avoid air bubbles adhesion during monitoring and measurement. 2. Automatic cleaning brush: Do not force the rotating shaft of the cleaning brush. 3. Installation attention: The sensor is best installed horizontally to avoid foreign matter depositing in the measurement monitoring window; when the foreign matter such as tree branches in the monitoring measurement area is easily blocked to the shaft, a protective cover should be selected for protection installation as much as possible.
25 | The others | 1. Problem: The operation interface cannot be connected or the measurement result is not displayed. Reason: The cable wiring method is wrong. The sensor address is incorrect. Solution: Check the wiring connection. Check the address out. 2. Problem: The measured value is too high, too low or the value is continuously unstable Reason: Sensor window is attached to foreign objects. Sensor self-cleaning damage. Solution: Clean the surface of the sensor window. Contact after-sales.
26
27

## Ph
Sl.No | Parameter | Configuration
1 | Model NO: | BH1701D
2 | pH Range: | 0-14pH
3 | pH Zero: | 7.00±0.25
4 | Temperature Range: | 0-60 degree celsius
5 | Output Signal: | RS485 & 4-20mA
6 | Pressure Range: | 0—0.3MPa
7 | Temperature Sensor: | NTC10K
8 | Housing Materials: | PA + GF
9 | Membrane Resistance: | ＜500MΩ
10 | Reference System: | Ag / AgCl / KCl
11 | Liquid Junction: | Ceramic Cores
12 | Electrolyte Solution: | KCl
13 | Double Salt Bridge: | Yes
14 | Threaded connection: | NPT3/4’’
15 | Cable Length: | 10m or Customize
16 | Wire Connection: | Pin, BNC or Customize
17 | Waterproof grade: | IP68
18 | Maintenance | The normal service life of the electrode is one year, which will be shortened due to bad environment or improper maintenance. During measurement, it shall be cleaned in distilled water (or deionized water) first and drained with filter paper to prevent impurities from being brought into the tested solution. The electrode bulb and the liquid volume shall be completely immersed in the tested solution. When the electrode is not in use, clean it and insert it into a protective sleeve with 3.3m potassium chloride solution, or insert the electrode into a container with 3.3m potassium chloride solution. There is dirt on the glass ball at the top of the electrode, which can be cleaned with 0.1n hydrochloric acid and then intruded into the 3.3mkCL solution for internal activation. Do not touch it by hand. To ensure the long-term accurate measurement, the electrode should be cleaned, maintained and calibrated regularly. The metal electrode will oxidize after being used for a period of time, and the electrode head will become black. Fine sandpaper can be used to polish the electrode head and remove the oxidized part, so as to ensure the sensitivity of the electrode measurement. Check whether the wiring terminal is dry. If it is stained, wipe it with anhydrous alcohol and use it after blow-dry. Long-term immersion in distilled water or protein solution should be avoided, and contact with silicone grease should be avoided. If the electrode is used for a long time, its glass film may become translucent or with sediment. At this time, hydrochloric acid can be used for washing and water flushing. It is recommended that users clean the electrode regularly and coordinate with the instrument calibration. When you use the above methods to maintain and maintain the electrode, the calibration procedure and normal measurement still cannot be carried out, indicating that the electrode can no longer recover the response, please replace the electrode.

## TDS
Sl.No | Parameter | Configuration
1 | Model NO: | BH3701D
2 | Power: | 9~36VDC
3 | Outlet: | MODBUS RS485 RTU /4-20mA
4 | Size Diameter: | 30mm*
5 | Size Length: | 165mm
6 | Weight: | 0.55KG (including 10m cable)
7 | Cable length: | Standard 10m cable, can be extended to 100m
8 | Material Main： | PP
9 | Material Cable： | PVC
10 | Measure material: | Graphite (2 Electrode)
11 | Waterproof grade: | IP68/NEMA6P
12 | Measurement range: | Con:1-30000us/cm;
13 |  | TDS：0-15000mg/L;
14 |  | Salinity：0-18.5ppt;；0-1.85%；0-18.5g/L
15 | Scope of Measurement: | 0～30000μS·cm-1 ；0~500000μS·cm-1
16 |  | Temperature：0-50 degree celsius
17 | Storage Temperature: | 0 to 45 degree celsius
18 | Accuracy Display: | ±1% F.S
19 | Accuracy Temperature： | ±0.5d egree celsius
20 | Pressure resistance: | ≤0.6Mpa
21 | Pressure Range: | ≤0.3Mpa
22 | Temperature compensation: | NTC10K
23 | Calibration: | Sample calibration, standard liquid calibration, Calibration of standard fluids and Alignment
24 | Connection methods: | 4 core cable
25 | Installation thread: | NPT3/4’’
26 | Application: | General application, river, lake , drink water,etc.
27 | Maintenance: | Check the instrument once a year for the best, if the user unconditionally can be sent to our company to help check. The maintenance period of the cleaning electrode can be determined according to the cleanliness of the water sample, and the surface inside and outside the electrode should be free of fouling deposition. For stainless steel conductance electrode and platinum conductance electrode, if contamination is found, it should be cleaned in time, brushed with 50% warm detergent (or household soap powder) and nylon brush, and then washed inside the electrode with distilled water. Ensure no grease deposition on the surface of internal and external electrodes. Do not touch the electrode with your fingers; clean the deposit with strong adhesion with 2% dilute hydrochloric acid solution, then wash it with clean water; for platinum black conductance electrode, do not brush the electrode head with brush or other substances, just rinse the electrode head with clean water; keep the cable and plug dry to ensure reliable contact.